###### Leishmania (IgG & IgM)

Catalog No. PI-301

##### For Research Use Only

# NAME AND INTENDED USE

The UBI MAGIWELTM Leishmania  donovani (PI-301) qualitative is a solid phase enzyme-linked immunosorbent assay (ELISA).  This test provides an easy method for detecting Leishmania antibodies (IgG & IgM) in human serum or plasma.  (For Professional Use Only)

SUMMARY AND EXPLANATION OF TEST

Leishmaniasis is caused by the intracellular protozoa parasites belonging to the genus Leishmania. Most forms of Leishmaniases are zoonotic; human are infected secondarily. These parasites of canines and rodents are transmitted from animal to human and occasionally from person to person  by the bite of Phlebotomus  sandflies. It is estimated that worldwide over 12 million people are infected with these parasites. The Leishmaniases are divided into the following three categories: Visceral Leishmaniais caused by Leishmania donovani, Cutaneous Leishmaniasis caused by L. tropica and L. brazilensis. Leishmania donovani is the etiological agent of Kala-azar. and occurs in the reticuloendothelial cells of viscera mainly the spleen, bone marrow and liver.  Serodiagnosisof visceral leishmaniais is more reliable than serodiagnosis of cutaneous disease. The direct agglutination test is positive early in the disease,complement fixation, complement fixation and  immunofluorescence are also widely used.

# PRINCIPLE OF THE ASSAY

UBI MAGIWELTM Leishmania donovani IgG & IgM is a solid phase enzyme-linked immunosorbent system employing plastic wells coated with purified Leishmania donovani antigens.  Incubation of diluted serum samples in the coated wells results in the binding of anti- Leishmania  donovani antibodies to the immobilized antigens.  Subsequent addition of the Antigen-enzyme conjugate (peroxidase) is in direct proportion to the amount of Leishmania antibody present in the serum sample.  Unbound enzyme conjugate is washed from the wells and a substrate and chromogen solution is added.  The intensity of the color formed as a result of enzyme activity is a direct measure of the anti-Leishmania antibody present in the serum samples and may be quantified by use of a photometric well reader at 450 nm wavelength.

# WARNING AND PRECAUTION

- UBI MAGIWELTM Leishmania IgG & IgM qualitative is designed for in vitro use only.

- The components in this kit are intended for use as an integral unit.  The components from different lots should not be mixed and used.

- References contains human serum should be treated as potentially infectious.  The references were found negative for Hepatitis B, C and HIV antibodies. However, because no test methods offer complete assurance of the absence of the HIV I/II, Hepatitis B, Hepatitis C virus or other infectious agents, these materials should be handled at the Biosafety Level 2 as recommended for any potentially infectious serum or blood specimen in the Center for Disease Control/National Institutes of Health Manual” Biosafety in Microbiological and Biomedical Laboratories”, 1984. Never pipet by mouth. Avoid contact with skin.

# MATERIALS PROVIDED

- Microwell Strips (96 wells):  Highly purified  Leishmania Antigens coated wells.  8 x 12 strips.

- Washing Buffer Concentrate (20X)  (50 mL):  Prepare working solution by adding purified water to 1 liter.

- Enzyme Conjugate (11 mL):  Leishmania Antigens conjugated to horseradish peroxidase.

- Substrate Solution (11 mL):  Buffer solution containing peroxide.

- Chromogen Solution (11 mL):  Tetramethylbenzidine solution.

- Reference Control Set (1.0 mL each vial): Negative, Positive Control. Ready to Use, DO NOT DILUTE

- Sample Diluent (11 mL)

- 2 N HCl solutions

- Well holder with wells.

MATERIALS REQUIRED BUT NOT PROVIDED

- Micro-well reader with wavelength at 450 nm.

- Pipetor with tips for measuring 5 uL, 50 uL  and 100 uL.

- Clean plastic washing bottle of 1000mL capacity for use in washing micro-wells with working washing buffer during testing procedure.

# REAGENT PREPARATION

Prepare the working washing buffer by adding the entire contents of the Wash Buffer Concentrate to 1000 mL distilled water in a clean plastic wash bottle.  Mix gently to dissolve.  Store at room temperature.

# STORAGE AND STABILITY

- Store the kits at 2-8oC and keep micro-wells in a dry bag with desiccants.

- Unopened reagents are stable until expiration of the kit.  Solution A and Solution B should be colorless; if the solution turns blue, it must be replaced.  Do not expose these reagents to strong light during storage or usage.

# PRECAUTIONS

This kit is designed for in vitro diagnostic use only.  The components of this kit are carefully matched and intended for use as an integral unit.  Components of different lots should not be used interchangeably.  Although all human materials used in the manufacture of this kit have been found negative for Hepatitis B antigen and for antibodies to HIV and HCV by required test methods, no test can offer complete assurance that infectious agents are not present, and therefore all calibrators, controls and samples should be handled as potentially infectious agents.

SPECIMEN COLLECTION AND HANDLING

Collect blood by venipuncture and allow clotting.  Separate the serum by centrifugation at room temperature.  Do not heat and inactivate serum.  If sera cannot be immediately assayed, they may be stored at –20o C for at least six months.  Avoid repeated freezing and thawing of samples.  Specimens obviously contaminated with bacteria should not be use.  Specimens turbid with high lipid concentrations should be clarified prior to assay.

# PREPARATION FOR ASSAY

- Bring all reagents and samples to room temperature (20-25oC) and mix gently before beginning the test.

- Have all reagents and samples ready before the start of the assay.  Once the test has begun it must be performed without any interruptions to get the most reliable and consistent results.

- Use new disposable tips for each specimen.

## ASSAY PROCEDURE  (30/30/15)

- Secure the desired number of coated wells in the holder.  Mark data sheet with sample identification.

- Well #1 is the blank well.  Dispense 100 uL of Negative Control to Wells #2 and #3. Add 100uL of Positive control to Wells # 4 and # 5. (DO NOT DILUTE NEGATIVE AND POSITIVE CONTROL).  Add 100uL of sample diluent to the remaining wells (Example: Wells #6, #7 for Patient 1, #8, #9 for Patient 2, etc.).  Add 5uL of Patient sample 1 to Wells #6, #7.  Add 5ul of Patient Sample 2 to Wells #8, #9, etc. mix well.

- Incubate 30 minutes at room temperature

- Wash five times with the washing buffer.

- Dispense 100 uL Enzyme Conjugate into each well except blank well.

- Incubate for 30 minutes at room temperature.

- Wash Five times with the Washing buffer.

- Dispense 100 uL of Solution A and 100 uL of Solution B.

- Incubate for 15 minutes at room temperature.

- Stop reaction by adding 50 uL of Stop Solution in each well.

- Zero a microreader on the blank and measures absorbance of each well at 450 nm.

** Dilution of Patients samples 1:21 out side of glass tube, then transfer 100 uL of diluted samples into each appropriate wells will produce better O.D. readings.

# INTERPRETATION OF RESULTS

- Negative control: O.D should be less than 0.2. Calculate the mean absorbance of the replicates of the Negative Control.

Absorbance Well #2 = 0.10

Well #3 = 0.13

Mean Negative Control= 0.12

- Cut-Off Value: Calculated by adding the mean of the Negative Control to the Factor of 0.2. For example: Cut-Off Value = 0.2 +0.12= 0.32

- Negative: Samples that developed no color or less intensity than the Cut –Off Value are considered Negative

- Positive: Samples that developed the color equal to or stronger than the Cut-Off Value are considered positive.

Specimens yielding absorbance reading within 10% of Cut-Off Value (gray zone) a new sample after one week should be retested with the old sample. If O. D. is less than Cut Off Value, the sample is considered Negative antibodies against Leishmania parasites.

# VALIDATION OF TEST

- Negative Control: mean absorbance value should be <0.2 units.

- Positive Control mean absorbance value should be greater than O.D. 0.5

- A test may be validated if the above criteria are met.

# LIMITATIONS OF THE PROCEDURE

The results obtained by means of this kit should be used as an aid for diagnosis and should not be interrupted as diagnostic by itself. Should negative results be obtained and other clinical findings suggest infection by Leishmania  parasites, a second serum should be obtained one week after the first and testing repeated. Initial testing may have occurred prior to significant antibody production in response to infection.

False-positive results may occur in  other Leishmania IgG antibodies.

# QUALITY CONTROL

Each laboratory should utilize internal controls several levels to monitor assay performance. The controls should be treated as unknown. Results obtained should be in agreement with the assigned values of the control.

# PERFORMANCE CHARACTERISTICS

Assay reproducibility was determined by assaying three specimens (S1, S2 and S3) in duplicate in 5 separate assays. The Intra-assay coefficient of variations (CV’s) were 4.56%, 7.62% and 8.21% for the S1, S2 and S3 respectively. The Inter assay CV’s were

5.53%, 7.47% and 7.94% for the S1, S2 and S3 respectively.

The UBI MAGIWELTM Leishmania ELISA is tested for the established reactivity.

Sensitivity of UBI MAGIWELTM is 100%: For assay of non-Leishmania 250 samples, the reactivities of 250 negative samples out of 250 were correctly identified, 6 of Leishmania positive samples out of 6 were identified as positive.

Specificity of UBI MAGIWELTM is 100% inclusive of Leishmania samples.

#### SPECIFICITY

UBI MAGIWELTM Leishmania (PI-301) IgG & IgM is specific. No cross reactivities found with the following tests: Toxoplasma IgG, IgM; CMV IgG,IgM; Rubella IgG, IgM, HSV I,II IgG, IgM, HCV, HBsAg, HIV, Cysticercosis, Chlamydia Trachomatis, Rubeola, RF factor, ANA, ENA, H.pyloi  Mono, DNA, Trypanosoma cruzi. etc. It might cross react with other Leishmania antibodies and other strains of Tryanomomiasis.

# REFERENCES

- Howard JG: Host immunity to Leishmania. In: Leishmaniasis( eds. Chang KP, Bray RS), pp. 140-162. Amsterdam: Elservier, 1985

- Kagan , IG: “Serodiagnosis of Parasitic Diseases” Manual of Clinical Laboratory Immunology, 3rd ed, Rose NR, Friedman H, and Fahey JL, eds, Washington, DC: American Society for Microbiology, 1986, 468

- Craig, CF Laboratory diagnosis of protozoan disease. Lea & Febiger, Philadelphia 1942.

- Kager PA et al: clinical, hematological and parasitological response to treatment of visceral Leishmaniasis in Kenya: A study of 64 patients. Trop Gfeog Med 36:21, 1984

- SWHO EXPERT COMMITTEE: The Leishmaniases. Geneva, WHO Tech Rep Ser 701, 1983

- Roitt IM, Delves PJ eds. : Encyclopedia of Immunology  Vol. II 959-962 Academic Press , 1992

September 22,2005

110 Pioneer Way, Suite C

Mountain View, CA 94041

Tel: (650) 961-2910

Fax: (650) 961-0766

www.unitedbiotech.com